{"text_format":"1","workbook":"xlsx","n_sheets":12,"sheets":["MECSB","MEBFP","MEFL","MEPE (488nm Ex)","MEPE (561nm Ex)","MEPTR (488nm Ex)","MEPTR (561nm Ex)","MECY (488nm Ex)","MECY (561nm Ex)","MEPCY7","MEAP","MEAPCY7"],"not_rendered":{"charts":24,"images":5,"macros":1,"external_links":0,"defined_names":12,"comments":12,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"MECSB","index":0,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[4,26,4,29],[5,26,5,29],[45,12,45,14],[46,12,46,14],[13,4,13,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE"]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MECSB","MECSB LOG","CALC.","RESIDUAL","CALC. MECSB",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MECSB","MECSB LOG","CALC.","RESIDUAL","CALC. MECSB",null,"CROSS CALIBRATION TABLE"]
[null,1,98.73798407593587,null,null,{"$f":"H$15*C6+H$16","v":1.8068635348670457},null,{"$f":"10^F6","v":64.100812570900374},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MECSB",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":0.49816472996259359},{"$f":"LOG10(U6)","v":-0.30262702373819428},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,121.66526366830416,198.01316823827568,{"$f":"LOG10(D7)","v":2.2966940726002023},{"$f":"H$15*C7+H$16","v":2.2966940726002023},{"$f":"((ABS(F7-E7))/F7)","v":0},{"$f":"10^F7","v":198.01316823827585},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":-0.30262702373819428},{"$f":"10^O7","v":0.49816472996259359},null,null,2,0,0.49816472996259359,-0.30262702373819428,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MECSB"]
[null,3,134.3291000301287,369.19089017029813,2.5672509762174123,2.5672509762174123,0,369.19089017029847,null,"1024 CH#","256 CH#",null,null,null,-0.30262702373819428,0.49816472996259359,null,null,3,0,0.49816472996259359,-0.30262702373819428,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,151.59614525829207,863.28316300774406,2.9361532707569413,2.9361532707569418,1.5124864708972642E-16,863.28316300774486,null,null,{"$f":"J9/4","v":0},null,null,null,-0.30262702373819428,0.49816472996259359,null,null,4,0,0.49816472996259359,-0.30262702373819428,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,174.88947494228321,2715.2154514476038,3.4338042964156079,3.4338042964156075,1.2932863131239808E-16,2715.2154514476038,null,null,0,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,5,0,0.49816472996259359,-0.30262702373819428,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,195.8391626760509,7609.9975131073916,3.8813845148462742,3.8813845148462733,2.2883030946891449E-16,7609.9975131073916,null,null,0,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,6,0,0.49816472996259359,-0.30262702373819428,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,216.05061994000579,20568.034245915442,4.3131927867386413,4.313192786738643,4.1184267136443405E-16,20568.034245915551,null,null,0,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,7,0,0.49816472996259359,-0.30262702373819428,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,227.82552928488985,36707.785309891769,4.5647581830137245,4.5647581830137245,0,36707.785309891769,null,null,0,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,8,0,0.49816472996259359,-0.30262702373819428,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":1.3160717989078186E-16},null,null,null,0,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X11)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.1364529348532231E-2},null,null,0,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":-0.30262702373819428},null,null,0,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":0.99999999999999956},null,null,null,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,-0.30262702373819428,0.49816472996259359,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MECSB/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MECSB","MECSB/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":0.49816472996259359},{"$f":"O39/N39","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":0.49816472996259359},{"$f":"O40/N40","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":0.49816472996259359},{"$f":"O41/N41","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":0.49816472996259359},{"$f":"O42/N42","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":0.49816472996259359},{"$f":"O43/N43","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MECSB "]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MECSB"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N54","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P39*N55","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P40*N56","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N57","v":0.49816472996259359}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]

{"sheet":"MEBFP","index":1,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[13,4,13,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[4,26,4,29],[5,26,5,29],[45,12,45,14],[46,12,46,14],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE"]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEBFP","MEBFP LOG","CALC.","RESIDUAL","CALC. MEBFP",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEBFP","MEBFP LOG","CALC.","RESIDUAL","CALC. MEBFP",null,"CROSS CALIBRATION TABLE"]
[null,1,63.423027465063782,null,null,{"$f":"H$15*C6+H$16","v":2.1114628525897379},null,{"$f":"10^F6","v":129.25961337733438},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEBFP",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":4.8715525347903093},{"$f":"LOG10(U6)","v":0.68766739033990876},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,96.44600481457293,712.52790410491571,{"$f":"LOG10(D7)","v":2.8528018769071992},{"$f":"H$15*C7+H$16","v":2.8528018769072006},{"$f":"((ABS(F7-E7))/F7)","v":4.6700320843680004E-16},{"$f":"10^F7","v":712.52790410491832},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":0.68766739033990865},{"$f":"10^O7","v":4.8715525347903093},null,null,2,0,4.8715525347903093,0.68766739033990876,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEBFP"]
[null,3,111.93407601185287,1586.7061234842738,3.2004964977897585,3.2004964977897594,2.7751269851833774E-16,1586.7061234842795,null,"1024 CH#","256 CH#",null,null,null,0.68766739033990865,4.8715525347903093,null,null,3,0,4.8715525347903093,0.68766739033990876,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,130.59405302619774,4162.9177323400882,3.6193978284357158,3.6193978284357171,3.6809096228197335E-16,4162.9177323401027,null,null,{"$f":"J9/4","v":0},null,null,null,0.68766739033990865,4.8715525347903093,null,null,4,0,4.8715525347903093,0.68766739033990876,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,154.95929290496974,14668.228856242078,4.1663776772790015,4.1663776772790033,4.2635521236767994E-16,14668.228856242156,null,null,0,null,null,null,0.68766739033990865,4.8715525347903093,null,null,5,0,4.8715525347903093,0.68766739033990876,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,177.12781217592746,46136.318057657503,4.6640429323817179,4.6640429323817187,1.9043101287375417E-16,46136.318057657663,null,null,0,null,null,null,0.68766739033990865,4.8715525347903093,null,null,6,0,4.8715525347903093,0.68766739033990876,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,200.50651521549312,154481.26970848482,5.1888758303280609,5.1888758303280618,1.7116971936558579E-16,154481.26970848537,null,null,0,null,null,null,0.68766739033990865,4.8715525347903093,null,null,7,0,4.8715525347903093,0.68766739033990876,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,217.90384970011658,379692.36261905928,5.5794318615439122,5.5794318615439131,1.5918796783268769E-16,379692.36261905998,null,null,0,null,null,null,0.68766739033990865,4.8715525347903093,null,null,8,0,4.8715525347903093,0.68766739033990876,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":2.9425011166811697E-16},null,null,null,0,null,null,null,0.68766739033990865,4.8715525347903093,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X11)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.2449187923646804E-2},null,null,0,null,null,null,0.68766739033990865,4.8715525347903093,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":0.68766739033990865},null,null,0,null,null,null,0.68766739033990865,4.8715525347903093,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":1.0000000000000004},null,null,null,null,null,null,0.68766739033990865,4.8715525347903093,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,0.68766739033990865,4.8715525347903093,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEBFP/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEBFP","MEBFP/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":4.8715525347903093},{"$f":"O39/N39","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":4.8715525347903093},{"$f":"O40/N40","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":4.8715525347903093},{"$f":"O41/N41","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":4.8715525347903093},{"$f":"O42/N42","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":4.8715525347903093},{"$f":"O43/N43","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MEBFP "]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEBFP"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N54","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P39*N55","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P40*N56","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N57","v":4.8715525347903093}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]

{"sheet":"MEFL","index":2,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[4,26,4,29],[5,26,5,29],[45,12,45,14],[46,12,46,14],[13,4,13,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE"]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEFL","MEFL LOG","CALC.","RESIDUAL","CALC. MEFL",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEFL","MEFL LOG","CALC.","RESIDUAL","CALC. MEFL",null,"CROSS CALIBRATION TABLE"]
[null,1,58.132804853338158,null,null,{"$f":"H$15*C6+H$16","v":1.754507901928464},null,{"$f":"10^F6","v":56.820872994567146},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEFL",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":3.3069898510397411},{"$f":"LOG10(U6)","v":0.51943286209114437},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,107.5518922555886,637.46079358379041,{"$f":"LOG10(D7)","v":2.8044534790607805},{"$f":"H$15*C7+H$16","v":2.80445347906078},{"$f":"((ABS(F7-E7))/F7)","v":1.5835142681660365E-16},{"$f":"10^F7","v":637.46079358379041},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":0.51943286209114437},{"$f":"10^O7","v":3.3069898510397411},null,null,2,0,3.3069898510397411,0.51943286209114437,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEFL"]
[null,3,127.75813084009687,1712.9716400923094,3.2337501728584628,3.2337501728584623,1.3732947386517231E-16,1712.9716400923094,null,"1024 CH#","256 CH#",null,null,null,0.51943286209114437,3.3069898510397411,null,null,3,0,3.3069898510397411,0.51943286209114437,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,148.63944692383757,4757.6160921295614,3.6773893944485212,3.6773893944485208,1.2076208478777658E-16,4757.6160921295614,null,null,{"$f":"J9/4","v":0},null,null,null,0.51943286209114437,3.3069898510397411,null,null,4,0,3.3069898510397411,0.51943286209114437,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,173.55437590133275,16096.288491869942,4.2067257472540884,4.2067257472540884,0,16096.288491869942,null,null,0,null,null,null,0.51943286209114437,3.3069898510397411,null,null,5,0,3.3069898510397411,0.51943286209114437,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,195.2415353537877,46503.508043595211,4.6674857156105967,4.6674857156105967,0,46503.508043595291,null,null,0,null,null,null,0.51943286209114437,3.3069898510397411,null,null,6,0,3.3069898510397411,0.51943286209114437,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,216.64580322602788,132505.96832340464,5.1222354401712824,5.1222354401712833,1.7339664099282896E-16,132505.96832340487,null,null,0,null,null,null,0.51943286209114437,3.3069898510397411,null,null,7,{"$f":"M56","v":0},{"$f":"O56","v":3.3069898510397411},0.51943286209114437,{"$error":"#DIV/0!"},{"$f":"((ABS(W12-V12))/W12)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W12","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,232.12306491281186,282527.98556667869,5.4510614729526639,5.451061472952663,1.6293678288295433E-16,282527.98556667869,null,null,0,null,null,null,0.51943286209114437,3.3069898510397411,null,null,8,{"$f":"M57","v":0},{"$f":"O57","v":3.3069898510397411},0.51943286209114437,{"$error":"#DIV/0!"},{"$f":"((ABS(W13-V13))/W13)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W13","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":1.0753948704933369E-16},null,null,null,0,null,null,null,0.51943286209114437,3.3069898510397411,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X11)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.1245750019343819E-2},null,null,0,null,null,null,0.51943286209114437,3.3069898510397411,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":0.51943286209114437},null,null,0,null,null,null,0.51943286209114437,3.3069898510397411,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":1.0000000000000004},null,null,null,null,null,null,0.51943286209114437,3.3069898510397411,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V7:V11,T7:T11)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,0.51943286209114437,3.3069898510397411,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEFL/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEFL","MEFL/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":3.3069898510397411},{"$f":"O39/N39","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":3.3069898510397411},{"$f":"O40/N40","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":3.3069898510397411},{"$f":"O41/N41","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":3.3069898510397411},{"$f":"O42/N42","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":3.3069898510397411},{"$f":"O43/N43","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MEFL "]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEFL"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N54","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J55)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N55","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P40*N56","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N57","v":3.3069898510397411}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]

{"sheet":"MEPE (488nm Ex)","index":3,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[13,4,13,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[4,26,4,29],[5,26,5,29],[45,12,45,14],[46,12,46,14],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE "]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEPE","MEPE LOG","CALC.","RESIDUAL","CALC. MEPE",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEPE","MEPE LOG","CALC.","RESIDUAL","CALC. MEPE",null,"CROSS CALIBRATION TABLE"]
[null,1,100.79947152000483,null,null,{"$f":"H$15*C6+H$16","v":2.2454167015014983},null,{"$f":"10^F6","v":175.96111347078264},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEPE",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":0.91236538688946034},{"$f":"LOG10(U6)","v":-3.9831199248141061E-2},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,120.71594714770563,497.67712684249881,{"$f":"LOG10(D7)","v":2.6969476811009483},{"$f":"H$15*C7+H$16","v":2.6969476811009483},{"$f":"((ABS(F7-E7))/F7)","v":0},{"$f":"10^F7","v":497.67712684249881},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":-3.9831199248141047E-2},{"$f":"10^O7","v":0.91236538688946034},null,null,2,{"$f":"M51","v":0},{"$f":"O51","v":0.91236538688946034},-3.9831199248141061E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEPE"]
[null,3,136.50563555785385,1134.7999884523322,3.054919322704456,3.0549193227044555,1.4536855574206124E-16,1134.7999884523322,null,"1024 CH#","256 CH#",null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,3,{"$f":"M52","v":0},{"$f":"O52","v":0.91236538688946034},-3.9831199248141061E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,155.29049937457202,3025.4868650437161,3.4807952718110293,3.4807952718110289,1.2758268590126148E-16,3025.4868650437161,null,null,{"$f":"J9/4","v":0},null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,4,{"$f":"M53","v":0},{"$f":"O53","v":0.91236538688946034},-3.9831199248141061E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,179.62722919333117,10778.01655569079,4.0325388463633312,4.0325388463633312,0,{"$f":"10^F10","v":10778.01655569079},null,null,0,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,5,{"$f":"M52","v":0},{"$f":"O52","v":0.91236538688946034},-3.9831199248141061E-2,{"$error":"#DIV/0!"},{"$f":"((ABS(W10-V10))/W10)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W10","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,201.19451199797535,33227.345972995849,4.521495653487122,4.5214956534871211,1.9643465078090561E-16,{"$f":"10^F11","v":33227.345972995849},null,null,0,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,6,{"$f":"M53","v":0},{"$f":"O53","v":0.91236538688946034},-3.9831199248141061E-2,{"$error":"#DIV/0!"},{"$f":"((ABS(W11-V11))/W11)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W11","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,224.01191553250698,109343.8179230306,5.0387942339414584,5.0387942339414575,1.762680471683743E-16,109343.8179230306,null,null,0,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,7,{"$f":"M54","v":0},{"$f":"O54","v":0.91236538688946034},-3.9831199248141061E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,242.24088688357128,283183.15608071786,5.4520674176662105,5.4520674176662105,0,283183.15608071786,null,null,0,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,8,{"$f":"M55","v":0},{"$f":"O55","v":0.91236538688946034},-3.9831199248141061E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":9.2236277084657511E-17},null,null,null,0,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.2671228988498271E-2},null,null,0,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":-3.9831199248141047E-2},null,null,0,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":1},null,null,null,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,-3.9831199248141047E-2,0.91236538688946034,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,21,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,56.5,null,"  CONVERSION",null,null,"THE CH# COLUMN BY",null,null,null,null,null,null,null,null,null,null,null,21]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS",null,null,null,null,null,null,null,null,null,null,null,56.5]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEPE/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEPE","MEPE/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":0.91236538688946034},{"$f":"O39/N39","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":0.91236538688946034},{"$f":"O40/N40","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":0.91236538688946034},{"$f":"O41/N41","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":0.91236538688946034},{"$f":"O42/N42","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":0.91236538688946034},{"$f":"O43/N43","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MEPE"]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEPE"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N54","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J55)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N55","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N56","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N57","v":0.91236538688946034}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]

{"sheet":"MEPE (561nm Ex)","index":4,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[3,12,3,15],[4,12,4,15],[4,26,4,29],[5,26,5,29],[13,4,13,5],[13,21,13,22],[46,12,46,14],[47,12,47,14],[33,12,33,15],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14],[45,12,45,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE "]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEPE","MEPE LOG","CALC.","RESIDUAL","CALC. MEPE",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEPE","MEPE LOG","CALC.","RESIDUAL","CALC. MEPE",null,"CROSS CALIBRATION TABLE"]
[null,1,100.79947152000483,null,null,{"$f":"H$15*C6+H$16","v":2.2979801102907356},null,{"$f":"10^F6","v":198.60039607655514},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEPE",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":1.1162671136231392},{"$f":"LOG10(U6)","v":4.7768130167708166E-2},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,120.71594714770563,552.82606764821617,{"$f":"LOG10(D7)","v":2.7425885133257286},{"$f":"H$15*C7+H$16","v":2.7425885133257291},{"$f":"((ABS(F7-E7))/F7)","v":1.6192338285248241E-16},{"$f":"10^F7","v":552.82606764821708},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":4.77681301677082E-2},{"$f":"10^O7","v":1.1162671136231392},null,null,2,{"$f":"M51","v":0},{"$f":"O51","v":1.1162671136231392},4.7768130167708166E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEPE"]
[null,3,136.50563555785385,1244.7208624710388,3.095071968716403,3.0950719687164039,2.8696535288272243E-16,1244.7208624710411,null,"1024 CH#","256 CH#",null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,3,{"$f":"M52","v":0},{"$f":"O52","v":1.1162671136231392},4.7768130167708166E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,155.29049937457202,3269.0282025788897,3.5144186673443802,3.5144186673443807,1.2636206777994158E-16,3269.0282025788952,null,null,{"$f":"J9/4","v":0},null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,4,{"$f":"M53","v":0},{"$f":"O53","v":1.1162671136231392},4.7768130167708166E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,179.61751637446557,11415.276780044425,4.0574864460574167,4.0574864460574176,2.188986781614887E-16,{"$f":"10^F10","v":11415.276780044465},null,null,0,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,5,{"$f":"M52","v":0},{"$f":"O52","v":1.1162671136231392},4.7768130167708166E-2,{"$error":"#DIV/0!"},{"$f":"((ABS(W10-V10))/W10)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W10","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,201.19451199797535,34606.980955277446,4.5391637139530925,4.5391637139530925,0,{"$f":"10^F11","v":34606.980955277446},null,null,0,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,6,{"$f":"M53","v":0},{"$f":"O53","v":1.1162671136231392},4.7768130167708166E-2,{"$error":"#DIV/0!"},{"$f":"((ABS(W11-V11))/W11)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W11","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,224.01191553250698,111823.07023615787,5.0485314121608624,5.0485314121608624,0,111823.07023615787,null,null,0,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,7,{"$f":"M54","v":0},{"$f":"O54","v":1.1162671136231392},4.7768130167708166E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,242.24088688357128,285409.59069159342,5.4554685627325661,5.4554685627325661,0,285409.59069159342,null,null,0,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,8,{"$f":"M55","v":0},{"$f":"O55","v":1.1162671136231392},4.7768130167708166E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":1.13449925953805E-16},null,null,null,0,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.2323648588538945E-2},null,null,0,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":4.77681301677082E-2},null,null,0,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":0.99999999999999933},null,null,null,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,4.77681301677082E-2,1.1162671136231392,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,21,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,56.5,null,"  CONVERSION",null,null,"THE CH# COLUMN BY",null,null,null,null,null,null,null,null,null,null,null,21]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS",null,null,null,null,null,null,null,null,null,null,null,56.5]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEPE/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEPE","MEPE/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":1.1162671136231392},{"$f":"O39/N39","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":1.1162671136231392},{"$f":"O40/N40","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":1.1162671136231392},{"$f":"O41/N41","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":1.1162671136231392},{"$f":"O42/N42","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":1.1162671136231392},{"$f":"O43/N43","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MEPE"]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEPE"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N54","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J55)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N55","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N56","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N57","v":1.1162671136231392}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]

{"sheet":"MEPTR (488nm Ex)","index":5,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[14,4,14,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[4,26,4,29],[5,26,5,29],[45,12,45,14],[46,12,46,14],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE"]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"  CROSS CALIBRATION TABLE"]
[null,null,null,null,null,null,null,null,null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEPTR","MEPTR LOG","CALC.","RESIDUAL","CALC. MEPTR",null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEPTR","MEPTR LOG","CALC.","RESIDUAL","CALC. MEPTR",null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEPTR",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":0.42720471914236335},{"$f":"LOG10(U6)","v":-0.36936395849315229},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,1,36.763911121367556,null,null,{"$f":"H$16*C7+H$17","v":0.46618669964377291},null,{"$f":"10^F7","v":2.9254097172617164},null,"  CONVERSION",null,null,null,null,{"$f":"H$16*N7+H$17","v":-0.36936395849315229},{"$f":"10^O7","v":0.42720471914236335},null,null,2,{"$f":"M51","v":0},{"$f":"O51","v":0.42720471914236335},-0.36936395849315229,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEPTR"]
[null,2,112.79270880089437,156.36115081044511,{"$f":"LOG10(D8)","v":2.194128858158181},{"$f":"H$16*C8+H$17","v":2.1941288581581824},{"$f":"((ABS(F8-E8))/F8)","v":6.0719662138190613E-16},{"$f":"10^F8","v":156.36115081044568},null,"1024 CH#","256 CH#",null,null,null,-0.36936395849315229,0.42720471914236335,null,null,3,{"$f":"M52","v":0},{"$f":"O52","v":0.42720471914236335},-0.36936395849315229,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,3,134.53413857617488,487.81655405205873,2.6882565340086177,2.6882565340086191,4.9558798153968008E-16,487.81655405206044,null,null,{"$f":"J9/4","v":0},null,null,null,-0.36936395849315229,0.42720471914236335,null,null,4,{"$f":"M53","v":0},{"$f":"O53","v":0.42720471914236335},-0.36936395849315229,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,125,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,4,155.78266968503252,1483.1390529883706,3.1711818705931138,3.1711818705931152,4.2011706799427751E-16,1483.1390529883761,null,null,0,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,5,{"$f":"M52","v":0},{"$f":"O52","v":0.42720471914236335},-0.36936395849315229,{"$error":"#DIV/0!"},{"$f":"((ABS(W10-V10))/W10)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W10","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,181.09478709833334,5577.7910333450682,3.7464622398036225,3.7464622398036234,2.3707123223179222E-16,5577.7910333450782,null,null,0,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,6,{"$f":"M53","v":0},{"$f":"O53","v":0.42720471914236335},-0.36936395849315229,{"$error":"#DIV/0!"},{"$f":"((ABS(W11-V11))/W11)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W11","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,203.18557251264841,17722.699964735217,4.2485298851843822,4.2485298851843831,2.0905547182271473E-16,17722.699964735279,null,null,0,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,7,{"$f":"M54","v":0},{"$f":"O54","v":0.42720471914236335},-0.36936395849315229,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,226.84392478243095,61125.772117794586,4.7862243580110988,4.7862243580111006,3.7113948417963623E-16,61125.772117794804,null,null,0,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,8,{"$f":"M55","v":0},{"$f":"O55","v":0.42720471914236335},-0.36936395849315229,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,246.60773648901318,171951.42532574356,5.2354057800913836,5.2354057800913845,1.696484393010358E-16,171951.42532574388,null,null,0,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G8:G14)","v":3.5854518549300612E-16},null,null,null,0,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E8:E14,C8:C14)","v":2.2727469212362846E-2},null,null,0,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E8:E14,C8:C14)","v":-0.36936395849315229},null,null,null,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E8:E14,C8:C14)","v":0.99999999999999978},null,null,null,null,null,null,-0.36936395849315229,0.42720471914236335,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEPTR/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEPTR","MEPTR/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":0.42720471914236335},{"$f":"O39/N39","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":0.42720471914236335},{"$f":"O40/N40","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":0.42720471914236335},{"$f":"O41/N41","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":0.42720471914236335},{"$f":"O42/N42","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":0.42720471914236335},{"$f":"O43/N43","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"Determination of New MEPTR"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,null,null,"values for RCP-30-5A"]
["Acceptable:"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEPTR"]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N54","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J55)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N55","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N56","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N57","v":0.42720471914236335}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]

{"sheet":"MEPTR (561nm Ex)","index":6,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[3,12,3,15],[4,12,4,15],[4,26,4,29],[5,26,5,29],[13,21,13,22],[14,4,14,5],[46,12,46,14],[47,12,47,14],[33,12,33,15],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14],[45,12,45,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE"]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"  CROSS CALIBRATION TABLE"]
[null,null,null,null,null,null,null,null,null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEPTR","MEPTR LOG","CALC.","RESIDUAL","CALC. MEPTR",null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEPTR","MEPTR LOG","CALC.","RESIDUAL","CALC. MEPTR",null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEPTR",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":0.54307109304457479},{"$f":"LOG10(U6)","v":-0.26514331351255338},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,1,36.763911121367556,null,null,{"$f":"H$16*C7+H$17","v":0.55680663779770279},null,{"$f":"10^F7","v":3.6041813727175041},null,"  CONVERSION",null,null,null,null,{"$f":"H$16*N7+H$17","v":-0.26514331351255338},{"$f":"10^O7","v":0.54307109304457479},null,null,2,{"$f":"M51","v":0},{"$f":"O51","v":0.54307109304457479},-0.26514331351255338,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEPTR"]
[null,2,112.79270880089437,180.56025264245145,{"$f":"LOG10(D8)","v":2.2566221537410933},{"$f":"H$16*C8+H$17","v":2.2566221537410946},{"$f":"((ABS(F8-E8))/F8)","v":5.9038134822062049E-16},{"$f":"10^F8","v":180.5602526424521},null,"1024 CH#","256 CH#",null,null,null,-0.26514331351255338,0.54307109304457479,null,null,3,{"$f":"M52","v":0},{"$f":"O52","v":0.54307109304457479},-0.26514331351255338,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,3,134.53413857617488,552.97646519027705,2.7427066480217017,2.7427066480217026,3.2383281687845212E-16,552.97646519027853,null,null,{"$f":"J9/4","v":0},null,null,null,-0.26514331351255338,0.54307109304457479,null,null,4,{"$f":"M53","v":0},{"$f":"O53","v":0.54307109304457479},-0.26514331351255338,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,125,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,4,155.78266968503252,1651.0915301008963,3.2177711495292418,3.2177711495292427,2.76022867515008E-16,1651.0915301009006,null,null,0,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,5,{"$f":"M52","v":0},{"$f":"O52","v":0.54307109304457479},-0.26514331351255338,{"$error":"#DIV/0!"},{"$f":"((ABS(W10-V10))/W10)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W10","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,181.09478709833334,6076.9739013203307,3.7836873711939671,3.783687371193968,2.3473884931985133E-16,6076.9739013203525,null,null,0,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,6,{"$f":"M53","v":0},{"$f":"O53","v":0.54307109304457479},-0.26514331351255338,{"$error":"#DIV/0!"},{"$f":"((ABS(W11-V11))/W11)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W11","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,203.18557251264841,18948.838392053447,4.2775825918563921,4.2775825918563921,0,18948.83839205348,null,null,0,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,7,{"$f":"M54","v":0},{"$f":"O54","v":0.54307109304457479},-0.26514331351255338,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,226.84392478243095,64050.824256392007,4.8065247229596961,4.806524722959697,1.8478598798368703E-16,64050.824256392116,null,null,0,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,8,{"$f":"M55","v":0},{"$f":"O55","v":0.54307109304457479},-0.26514331351255338,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,246.60773648901318,177171.7919700653,5.2483945777828591,5.2483945777828609,3.3845718211046876E-16,177171.79197006626,null,null,0,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G8:G14)","v":2.7831700743258395E-16},null,null,null,0,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E8:E14,C8:C14)","v":2.23575219893438E-2},null,null,0,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E8:E14,C8:C14)","v":-0.26514331351255338},null,null,null,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E8:E14,C8:C14)","v":1},null,null,null,null,null,null,-0.26514331351255338,0.54307109304457479,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEPTR/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEPTR","MEPTR/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":0.54307109304457479},{"$f":"O39/N39","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":0.54307109304457479},{"$f":"O40/N40","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":0.54307109304457479},{"$f":"O41/N41","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":0.54307109304457479},{"$f":"O42/N42","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":0.54307109304457479},{"$f":"O43/N43","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"Determination of New MEPTR"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,null,null,"values for RCP-30-5A"]
["Acceptable:"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEPTR"]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N54","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J55)*(256/LOG10(262144))","v":{"$error":"#NUM!"}},null,null,1,{"$f":"P39*N55","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N56","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,null,1,{"$f":"P41*N57","v":0.54307109304457479}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]

{"sheet":"MECY (488nm Ex)","index":7,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[13,4,13,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[4,26,4,29],[5,26,5,29],[45,12,45,14],[46,12,46,14],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE "]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MECY","MECY LOG","CALC.","RESIDUAL","CALC. MECY",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MECY","MECY LOG","CALC.","RESIDUAL","CALC. MECY",null,"CROSS CALIBRATION TABLE"]
[null,1,22.541688899145331,null,null,{"$f":"H$15*C6+H$16","v":1.3790273161469253},null,{"$f":"10^F6","v":23.93466295322218},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MECY",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":7.9088003984524899},{"$f":"LOG10(U6)","v":0.89811061499684941},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,103.87599286296324,1300.9415187430327,{"$f":"LOG10(D7)","v":3.114257774149559},{"$f":"H$15*C7+H$16","v":3.1142577741495594},{"$f":"((ABS(F7-E7))/F7)","v":1.42598732043411E-16},{"$f":"10^F7","v":1300.941518743035},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":0.89811061499684941},{"$f":"10^O7","v":7.9088003984524899},null,null,2,{"$f":"M51","v":0},{"$f":"O51","v":7.9088003984524899},0.89811061499684941,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MECY"]
[null,3,124.84376362005843,3644.1449745772557,3.5615956461293097,3.5615956461293101,1.2468827289046478E-16,3644.144974577262,null,"1024 CH#","256 CH#",null,null,null,0.89811061499684941,7.9088003984524899,null,null,3,{"$f":"M52","v":0},{"$f":"O52","v":7.9088003984524899},0.89811061499684941,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,146.18239229469128,10395.505282776794,4.0168456034499416,4.0168456034499407,2.2111340772901429E-16,10395.505282776794,null,null,{"$f":"J9/4","v":0},null,null,null,0.89811061499684941,7.9088003984524899,null,null,4,{"$f":"M53","v":0},{"$f":"O53","v":7.9088003984524899},0.89811061499684941,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,171.68778698359762,36390.801624036911,4.5609916224194675,4.5609916224194675,0,36390.801624036976,null,null,0,null,null,null,0.89811061499684941,7.9088003984524899,null,null,5,{"$f":"M52","v":0},{"$f":"O52","v":7.9088003984524899},0.89811061499684941,{"$error":"#DIV/0!"},{"$f":"((ABS(W10-V10))/W10)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W10","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,194.40349514031263,111076.18251885691,5.0456209454436811,5.045620945443682,1.7602955697696077E-16,111076.18251885709,null,null,0,null,null,null,0.89811061499684941,7.9088003984524899,null,null,6,{"$f":"M53","v":0},{"$f":"O53","v":7.9088003984524899},0.89811061499684941,{"$error":"#DIV/0!"},{"$f":"((ABS(W11-V11))/W11)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W11","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,218.55830217935917,363875.47881560004,5.5609527899256337,5.5609527899256328,1.5971695017968367E-16,363875.47881560004,null,null,0,null,null,null,0.89811061499684941,7.9088003984524899,null,null,7,{"$f":"M56","v":0},{"$f":"O56","v":7.9088003984524899},0.89811061499684941,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,237.16407096550742,907608.22295661154,5.9578984219512234,5.9578984219512225,1.4907579095805482E-16,907608.22295661154,null,null,0,null,null,null,0.89811061499684941,7.9088003984524899,null,null,8,{"$f":"M57","v":0},{"$f":"O57","v":7.9088003984524899},0.89811061499684941,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":1.3903181582536989E-16},null,null,null,0,null,null,null,0.89811061499684941,7.9088003984524899,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.133454610707141E-2},null,null,0,null,null,null,0.89811061499684941,7.9088003984524899,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":0.89811061499684941},null,null,0,null,null,null,0.89811061499684941,7.9088003984524899,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":1},null,null,null,null,null,null,0.89811061499684941,7.9088003984524899,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,0.89811061499684941,7.9088003984524899,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MECY/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MECY","MECY/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":7.9088003984524899},{"$f":"O39/N39","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":7.9088003984524899},{"$f":"O40/N40","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":7.9088003984524899},{"$f":"O41/N41","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":7.9088003984524899},{"$f":"O42/N42","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":7.9088003984524899},{"$f":"O43/N43","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MECY"]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MECY"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,1,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":0},null,null,1,{"$f":"P40*N54","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,119,{"$f":"LOG10(J55)*(256/LOG10(262144))","v":98.059630411490744},null,null,1,{"$f":"P41*N55","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,356,120.54376435152211,null,null,1,{"$f":"P39*N56","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,1196,145.4080238108175,null,null,1,{"$f":"P39*N57","v":7.9088003984524899}]
[null,null,null,null,null,null,null,null,null,3765,168.93773012706546]
[null,null,null,null,null,null,null,null,null,11452,191.76265891554638]
[null,null,null,null,null,null,null,null,null,39824,217.33476287973031]
[null,null,null,null,null,null,null,null,null,100094,236.24527602903785]

{"sheet":"MECY (561nm Ex)","index":8,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[3,12,3,15],[4,12,4,15],[4,26,4,29],[5,26,5,29],[13,4,13,5],[13,21,13,22],[46,12,46,14],[47,12,47,14],[33,12,33,15],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14],[45,12,45,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE "]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MECY","MECY LOG","CALC.","RESIDUAL","CALC. MECY",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MECY","MECY LOG","CALC.","RESIDUAL","CALC. MECY",null,"CROSS CALIBRATION TABLE"]
[null,1,22.541688899145331,null,null,{"$f":"H$15*C6+H$16","v":1.3367400994616161},null,{"$f":"10^F6","v":21.714013298049288},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MECY",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":7.1088898549366952},{"$f":"LOG10(U6)","v":0.85180178532631301},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,103.87599286296324,1220.3410804053776,{"$f":"LOG10(D7)","v":3.0864812311954335},{"$f":"H$15*C7+H$16","v":3.0864812311954331},{"$f":"((ABS(F7-E7))/F7)","v":1.4388203801843994E-16},{"$f":"10^F7","v":1220.3410804053776},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":0.8518017853263129},{"$f":"10^O7","v":7.1088898549366952},null,null,2,{"$f":"M51","v":0},{"$f":"O51","v":7.1088898549366952},0.85180178532631301,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MECY"]
[null,3,124.84376362005843,3447.9417239026498,3.5375599168970524,3.537559916897052,1.2553545954907603E-16,3447.9417239026498,null,"1024 CH#","256 CH#",null,null,null,0.8518017853263129,7.1088898549366952,null,null,3,{"$f":"M52","v":0},{"$f":"O52","v":7.1088898549366952},0.85180178532631301,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,146.18239229469128,9922.402768867616,3.9966168518824352,3.9966168518824352,0,9922.402768867616,null,null,{"$f":"J9/4","v":0},null,null,null,0.8518017853263129,7.1088898549366952,null,null,4,{"$f":"M53","v":0},{"$f":"O53","v":7.1088898549366952},0.85180178532631301,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,171.68778698359762,35100.494269389033,4.5453132320486409,4.54531323204864,1.9540532728034016E-16,35100.494269389033,null,null,0,null,null,null,0.8518017853263129,7.1088898549366952,null,null,5,{"$f":"M52","v":0},{"$f":"O52","v":7.1088898549366952},0.85180178532631301,{"$error":"#DIV/0!"},{"$f":"((ABS(W10-V10))/W10)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W10","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,194.40349514031263,108142.20348954746,5.0339952144499058,5.0339952144499041,3.5287217482870892E-16,108142.20348954709,null,null,0,null,null,null,0.8518017853263129,7.1088898549366952,null,null,6,{"$f":"M53","v":0},{"$f":"O53","v":7.1088898549366952},0.85180178532631301,{"$error":"#DIV/0!"},{"$f":"((ABS(W11-V11))/W11)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W11","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,218.55830217935917,357796.81044927321,5.5536364647655789,5.5536364647655763,4.7978208080511223E-16,357796.81044927199,null,null,0,null,null,null,0.8518017853263129,7.1088898549366952,null,null,7,{"$f":"M56","v":0},{"$f":"O56","v":7.1088898549366952},0.85180178532631301,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,237.16407096550742,899293.61712504807,5.9539015109524254,5.9539015109524254,0,899293.61712504807,null,null,0,null,null,null,0.8518017853263129,7.1088898549366952,null,null,8,{"$f":"M57","v":0},{"$f":"O57","v":7.1088898549366952},0.85180178532631301,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":1.8535386864023963E-16},null,null,null,0,null,null,null,0.8518017853263129,7.1088898549366952,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.15129539008805E-2},null,null,0,null,null,null,0.8518017853263129,7.1088898549366952,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":0.8518017853263129},null,null,0,null,null,null,0.8518017853263129,7.1088898549366952,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":1},null,null,null,null,null,null,0.8518017853263129,7.1088898549366952,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,0.8518017853263129,7.1088898549366952,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MECY/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MECY","MECY/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":7.1088898549366952},{"$f":"O39/N39","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":7.1088898549366952},{"$f":"O40/N40","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":7.1088898549366952},{"$f":"O41/N41","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":7.1088898549366952},{"$f":"O42/N42","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":7.1088898549366952},{"$f":"O43/N43","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MECY"]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MECY"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,1,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":0},null,null,1,{"$f":"P40*N54","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,119,{"$f":"LOG10(J55)*(256/LOG10(262144))","v":98.059630411490744},null,null,1,{"$f":"P41*N55","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,356,120.54376435152211,null,null,1,{"$f":"P39*N56","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,1196,145.4080238108175,null,null,1,{"$f":"P39*N57","v":7.1088898549366952}]
[null,null,null,null,null,null,null,null,null,3765,168.93773012706546]
[null,null,null,null,null,null,null,null,null,11452,191.76265891554638]
[null,null,null,null,null,null,null,null,null,39824,217.33476287973031]
[null,null,null,null,null,null,null,null,null,100094,236.24527602903785]

{"sheet":"MEPCY7","index":9,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[13,4,13,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[5,26,5,29],[4,26,4,29],[45,12,45,14],[46,12,46,14],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE"]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEPCY7","MEPCY7 LOG","CALC.","RESIDUAL","CALC. MEPCY7",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEPCY7","MEPCY7 LOG","CALC.","RESIDUAL","CALC. MEPCY7",null,"CROSS CALIBRATION TABLE"]
[null,1,49.200805242841568,null,null,{"$f":"H$15*C6+H$16","v":3.6816287651711077E-2},null,{"$f":"10^F6","v":1.0884695588183018},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEPCY7",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":3.8293253744838641E-2},{"$f":"LOG10(U6)","v":-1.4168777304558382},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,106.43649105442499,53,{"$f":"LOG10(D7)","v":1.7242758696007889},1.7279100345104075,{"$f":"((ABS(F7-E7))/F7)","v":2.103214193468273E-3},53.445363415081239,null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":-1.416877730455838},{"$f":"10^O7","v":3.8293253744838641E-2},null,null,2,0,3.8293253744838641E-2,-1.4168777304558382,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEPCY7"]
[null,3,124.70306595238294,186,2.2695129442179165,2.2676168651082351,8.3615496905861512E-4,185.18971576064257,null,"1024 CH#","256 CH#",null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,3,0,3.8293253744838641E-2,-1.4168777304558382,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,145.86607912572285,785,2.8948696567452528,2.8929022746329411,6.8007209561244113E-4,781.45194172606568,null,null,{"$f":"J9/4","v":0},null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,4,0,3.8293253744838641E-2,-1.4168777304558382,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,171.28347184295379,4416.2732254436105,3.6450559344294238,3.643888189732468,3.2046666531816172E-4,4404.4145608630206,null,null,0,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,5,0,3.8293253744838641E-2,-1.4168777304558382,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,194.32299695613656,21133.014222982642,4.3249614453271761,4.3246172890177981,7.9580755099874264E-5,21116.274016786832,null,null,0,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,6,0,3.8293253744838641E-2,-1.4168777304558382,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,219.15813941227842,114250.03325954376,5.0578563351701957,5.0583999547516072,1.0746868303697906E-4,114393.13303550699,null,null,0,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,7,0,3.8293253744838641E-2,-1.4168777304558382,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,237.45233294756846,396027.39750172046,5.5977252318149784,5.5989228095522847,2.1389431110983579E-4,397120.95988033031,null,null,0,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,8,0,3.8293253744838641E-2,-1.4168777304558382,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":6.201216675291685E-4},null,null,null,0,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.954614281072266E-2},null,null,0,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":-1.416877730455838},null,null,0,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":0.99999807622145331},null,null,null,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,-1.416877730455838,3.8293253744838641E-2,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEPCY7/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEPCY7","MEPCY7/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":3.8293253744838641E-2},{"$f":"O39/N39","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":3.8293253744838641E-2},{"$f":"O40/N40","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":3.8293253744838641E-2},{"$f":"O41/N41","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":3.8293253744838641E-2},{"$f":"O42/N42","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":3.8293253744838641E-2},{"$f":"O43/N43","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MEPCY7"]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEPCY7"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,8,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":42.666666666666664},null,null,1,{"$f":"P39*N54","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,146,102.25528261518247,null,null,1,{"$f":"P39*N55","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,377,121.71976125537657,null,null,1,{"$f":"P39*N56","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,1340,147.74069028046415,null,null,1,{"$f":"P39*N57","v":3.8293253744838641E-2}]
[null,null,null,null,null,null,null,null,null,4198,171.17136286282175]
[null,null,null,null,null,null,null,null,null,13333,194.88304086757577]
[null,null,null,null,null,null,null,null,null,50516,222.21443963046033]
[null,null,null,null,null,null,null,null,null,148865,244.38961881320574]

{"sheet":"MEAP","index":10,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[5,26,5,29],[4,26,4,29],[45,12,45,14],[46,12,46,14],[13,4,13,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE"]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEAP","MEAP LOG","CALC.","RESIDUAL","CALC. MEAP",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEAP","MEAP LOG","CALC.","RESIDUAL","CALC. MEAP",null,"CROSS CALIBRATION TABLE"]
[null,1,50.986133343589778,null,null,{"$f":"H$15*C6+H$16","v":1.0913611807538053},null,{"$f":"10^F6","v":12.341307706987749},null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEAP",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":0.97395802617693783},{"$f":"LOG10(U6)","v":-1.1459759130146614E-2},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,114.25312713297706,288.28027941824018,{"$f":"LOG10(D7)","v":2.4598149343314208},{"$f":"H$15*C7+H$16","v":2.4598149343314204},{"$f":"((ABS(F7-E7))/F7)","v":1.8053765088257194E-16},{"$f":"10^F7","v":288.2802794182399},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":-1.145975913014663E-2},{"$f":"10^O7","v":0.97395802617693783},null,null,2,0,0.97395802617693783,-1.1459759130146614E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEAP"]
[null,3,137.58851225807194,921.62689216871718,2.9645551385510185,2.964555138551018,1.4979961211553635E-16,921.62689216871718,null,"1024 CH#","256 CH#",null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,3,0,0.97395802617693783,-1.1459759130146614E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,160.97483107132763,2953.9084375101997,3.4703970293305808,3.4703970293305813,1.2796495792751551E-16,2953.9084375102052,null,null,{"$f":"J9/4","v":0},null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,4,0,0.97395802617693783,-1.1459759130146614E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,186.40910286353375,10484.211574363229,4.0205357765488206,4.0205357765488214,2.2091046294892735E-16,10484.211574363266,null,null,0,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,5,0,0.97395802617693783,-1.1459759130146614E-2,{"$error":"#DIV/0!"},{"$f":"((ABS(W10-V10))/W10)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W10","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,208.58392929260503,31635.398090562034,4.5001733038638125,4.5001733038638143,3.9473076245198915E-16,31635.398090562147,null,null,0,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,6,0,0.97395802617693783,-1.1459759130146614E-2,{"$error":"#DIV/0!"},{"$f":"((ABS(W11-V11))/W11)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W11","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,229.14107899395438,88068.535210801914,4.9448207730531495,4.9448207730531495,0,88068.53521080206,null,null,0,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,7,0,0.97395802617693783,-1.1459759130146614E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,238.44828005681964,140001.87963196187,5.1461338664518754,5.1461338664518754,0,140001.8796319621,null,null,0,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,8,0,0.97395802617693783,-1.1459759130146614E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G7:G13)","v":1.5342049233236289E-16},null,null,null,0,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E7:E13,C7:C13)","v":2.1629821042756205E-2},null,null,0,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E7:E13,C7:C13)","v":-1.145975913014663E-2},null,null,0,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E7:E13,C7:C13)","v":1.0000000000000004},null,null,null,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,-1.145975913014663E-2,0.97395802617693783,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEAP/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEAP","MEAP/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":0.97395802617693783},{"$f":"O39/N39","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":0.97395802617693783},{"$f":"O40/N40","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":0.97395802617693783},{"$f":"O41/N41","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":0.97395802617693783},{"$f":"O42/N42","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":0.97395802617693783},{"$f":"O43/N43","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MEAP"]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEAP"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,39,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":75.17016489048531},null,null,1,{"$f":"P39*N54","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,311,{"$f":"LOG10(J55)*(256/LOG10(262144))","v":117.77096206407968},null,null,1,{"$f":"P39*N55","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,746,135.72311922140781,null,null,1,{"$f":"P39*N56","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,2854,163.25361236587753,null,null,1,{"$f":"P39*N57","v":0.97395802617693783}]
[null,null,null,null,null,null,null,null,null,9044,186.91904393602175]
[null,null,null,null,null,null,null,null,null,25985,208.57445601500822]
[null,null,null,null,null,null,null,null,null,80493,231.77352114601644]
[null,null,null,null,null,null,null,null,null,169790,247.08824130067222]

{"sheet":"MEAPCY7","index":11,"n_rows":61,"n_cols":30,"rows_kept":61,"cut":false,"hidden":false,"merged":[[13,4,13,5],[3,12,3,15],[4,12,4,15],[33,12,33,15],[5,26,5,29],[4,26,4,29],[45,12,45,14],[46,12,46,14],[47,12,47,14],[13,21,13,22],[34,12,34,15],[35,12,35,15],[36,12,36,15],[44,12,44,14]]}
[null,"TYPE THE NAME OF YOUR LAB HERE"]
[]
[null,"PMT LINEARITY QC RECORD",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,"Normalization Graph for Instrument B"]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 1",null,null,"CROSS CALIBRATION TABLE"]
[null,"PEAK #","CH #","MEAPCY7","MEAPCY7 LOG","CALC.","RESIDUAL","CALC. MEAPCY7 ",null,"  1024 REL. CH#",null,null,"FOR UNKNOWN SAMPLES",null,null,null,null,null,"PEAK #","CH #","MEAPCY7","MEAPCY7 LOG","CALC.","RESIDUAL","CALC. MEAPCY7 ",null,"CROSS CALIBRATION TABLE"]
[null,1,14.222222222222223,null,null,{"$f":"H$15*C6+H$16","v":0.25028840329357521},null,null,null," to 256 REL. CH#",null,null,"Sample","CH#","Calc. ","Calc. MEAPCY7",null,null,1,{"$f":"M50","v":0},{"$f":"O50","v":0.96166071601590164},{"$f":"LOG10(U6)","v":-1.6978124586096009E-2},{"$f":"Y$15*T6+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"((ABS(W6-V6))/W6)*10","v":{"$error":"#DIV/0!"}},{"$f":"10^W6","v":{"$error":"#DIV/0!"}},null,"FOR NORMALIZATION GRAPH"]
[null,2,113.53590861787464,130.80033316967413,{"$f":"LOG10(D7)","v":2.1166088502081775},2.1166088502081775,{"$f":"((ABS(F7-E7))/F7)","v":0},{"$f":"10^F7","v":130.80033316967413},null,"  CONVERSION",null,null,null,null,{"$f":"H$15*N7+H$16","v":-1.6978124586096044E-2},{"$f":"10^O7","v":0.96166071601590164},null,null,2,0,0.96166071601590164,-1.6978124586096009E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,"Sample","CH#","Calc. ","Calc. MEAPCY7"]
[null,3,136.61126425687706,355.0136362364575,2.5502450348255445,2.5502450348255441,1.7413589823161507E-16,355.0136362364575,null,"1024 CH#","256 CH#",null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,3,0,0.96166071601590164,-1.6978124586096009E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,200,{"$f":"Y$15*AB8+Y$16","v":{"$error":"#DIV/0!"}},{"$f":"10^AC8","v":{"$error":"#DIV/0!"}}]
[null,4,160.19706338981095,985.08467845048085,2.9934735643078838,2.9934735643078834,1.4835247424432821E-16,985.08467845048085,null,null,{"$f":"J9/4","v":0},null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,4,0,0.96166071601590164,-1.6978124586096009E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,5,185.77612077335229,2979.6122912424098,3.4741597571204279,3.4741597571204275,1.2782636404094091E-16,2979.6122912424098,null,null,0,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,5,0,0.96166071601590164,-1.6978124586096009E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,6,208.29140718605939,7893.5255908279223,3.8972710212896864,3.8972710212896859,1.1394876246073963E-16,7893.5255908279223,null,null,0,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,6,0,0.96166071601590164,-1.6978124586096009E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,7,229.53175756221273,19788.982336274079,4.2964234608509333,4.2964234608509324,2.0672506511362736E-16,19788.982336274079,null,null,0,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,7,0,0.96166071601590164,-1.6978124586096009E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,8,239.93099309854674,31034.713708521653,4.4918477434258541,4.4918477434258541,0,31034.713708521653,null,null,0,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,8,0,0.96166071601590164,-1.6978124586096009E-2,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},{"$error":"#DIV/0!"},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(G8:G13)","v":1.2849809401520852E-16},null,null,null,0,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,null,null,null,"Ave Residual",null,{"$f":"AVERAGE(X6:X13)","v":{"$error":"#DIV/0!"}},null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(E8:E13,C8:C13)","v":1.8792177741539386E-2},null,null,0,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,null,null,null,null,null,"Slope:",{"$f":"SLOPE(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(E8:E13,C8:C13)","v":-1.6978124586096044E-2},null,null,0,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,null,null,null,null,null,"Intercept:",{"$f":"INTERCEPT(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(E8:E13,C8:C13)","v":0.99999999999999978},null,null,null,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,null,null,null,null,null,"Rsq:",{"$f":"RSQ(V6:V13,T6:T13)","v":{"$error":"#DIV/0!"}},null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,null,null,null,null,null,-1.6978124586096044E-2,0.96166071601590164,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 2",null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,null,{"$error":"#DIV/0!"},{"$error":"#DIV/0!"}]
[null,null,null,null,null,null,null,null,null,"  1024 MEAN CH#",null,null,"CONVERTED VALUES"]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"CAN BE COPIED TO"]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"THE CH# COLUMN BY"]
[null,null,null,null,null,null,null,null,null,"1024 CH#","256 CH#",null,"SELECTING THE CELLS"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J24*10)*(64)","v":{"$error":"#NUM!"}},null,"TO BE COPIED; THEN,"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"CLICKING [COPY],"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"[PASTE SPECIAL], AND"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"FINALLY [PASTE VALUE]"]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[]
[]
[null,null,null,null,null,null,null,null,null,"     TABLE NO. 3",null,null,"     TABLE NO. 5"]
[null,null,null,null,null,null,null,null,null,"  104 MEAN CH#",null,null," Determination of the "]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,"MEAPCY7/CH# for the "]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,"  normalization cell or particle"]
[null,null,null,null,null,null,null,null,null,"104 CH#","256 CH#",null,"256 CH#","104 CH#","Calc. MEAPCY7","MEAPCY7/CH#"]
[null,null,null,null,null,null,null,null,null,null,{"$f":"LOG10(J39)*(64)","v":{"$error":"#NUM!"}},null,{"$f":"N7","v":0},{"$f":"10^(4*(M39/256))","v":1},{"$f":"P7","v":0.96166071601590164},{"$f":"O39/N39","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N8","v":0},{"$f":"10^(4*(M40/256))","v":1},{"$f":"P8","v":0.96166071601590164},{"$f":"O40/N40","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N9","v":0},{"$f":"10^(4*(M41/256))","v":1},{"$f":"P9","v":0.96166071601590164},{"$f":"O41/N41","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N10","v":0},{"$f":"10^(4*(M42/256))","v":1},{"$f":"P10","v":0.96166071601590164},{"$f":"O42/N42","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,{"$f":"N11","v":0},{"$f":"10^(4*(M43/256))","v":1},{"$f":"P11","v":0.96166071601590164},{"$f":"O43/N43","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"}]
[null,null,null,null,null,null,null,null,null,null,{"$error":"#NUM!"},null,"     TABLE NO. 6"]
["Rainbow Calibration Particles (RCP-30-5A)",null,null,null,"Lot No.:",null,"File # ",null,null,null,{"$error":"#NUM!"},null,"Determination of New MEAPCY7"]
["Acceptable:",null,null,null,null,null,null,null,null,null,null,null,"values for RCP-30-5A"]
[]
["Action taken if not linear:",null,null,"By:",null,null,"Date:",null,null,"     TABLE NO. 4",null,null,"256 CH#","104 CH#","New MEAPCY7"]
[null,null,null,null,null,null,null,null,null,"  105 MEAN CH#",null,null,null,{"$f":"10^(4*(M50/256))","v":1},{"$f":"P39*N50","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null," to 256 REL. CH#",null,null,null,1,{"$f":"P39*N51","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,"  CONVERSION",null,null,null,1,{"$f":"P39*N52","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,"105 CH#","256 CH#",null,null,1,{"$f":"P39*N53","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,1,{"$f":"LOG10(J54)*(256/LOG10(262144))","v":0},null,null,1,{"$f":"P39*N54","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,228,111.40110242367633,null,null,1,{"$f":"P39*N55","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,606,131.45847443161529,null,null,1,{"$f":"P39*N56","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,2299,158.81657956576967,null,null,1,{"$f":"P39*N57","v":0.96166071601590164}]
[null,null,null,null,null,null,null,null,null,7224,182.30872430194998]
[null,null,null,null,null,null,null,null,null,21540,204.72505784080454]
[null,null,null,null,null,null,null,null,null,69266,228.69133855960027]
[null,null,null,null,null,null,null,null,null,158340,245.65569750662809]