## MECSB
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MECSB | MECSB LOG | CALC. | RESIDUAL | CALC. MECSB | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MECSB | MECSB LOG | CALC. | RESIDUAL | CALC. MECSB | CROSS CALIBRATION TABLE
| 1 | 115.7625 |  |  | 3.3232 |  | 2104.6093 | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MECSB |  | 1 | 0 | 7.2992 | 0.86327 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 129.8753 | 3827.0731 | 3.5829 | 3.6231 | 1.109% | 4198.1856 | CONVERSION |  |  |  | 0.86327 | 7.2992 |  | 2 | 0 | 7.2992 | 0.86327 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MECSB
| 3 | 139.136 | 6410.855 | 3.8069 | 3.8198 | 0.3385% | 6604.6017 | 1024 CH# | 256 CH# |  |  | 0.86327 | 7.2992 |  | 3 | 0 | 7.2992 | 0.86327 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 4 | 152.2824 | 12405.9172 | 4.0936 | 4.0992 | 0.136% | 12566.1677 |  | 0 |  |  | 0.86327 | 7.2992 |  | 4 | 0 | 7.2992 | 0.86327 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 173.9893 | 37097.3178 | 4.5693 | 4.5605 | 0.1947% | 36346.4906 |  | 0 |  |  | 0.86327 | 7.2992 |  | 5 | 0 | 7.2992 | 0.86327 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 193.2184 | 92487.771 | 4.9661 | 4.9691 | 0.06013% | 93126.248 |  | 0 |  |  | 0.86327 | 7.2992 |  | 6 | 0 | 7.2992 | 0.86327 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 214.5902 | 267142.3048 | 5.4267 | 5.4232 | 0.06507% | 264980.3173 |  | 0 |  |  | 0.86327 | 7.2992 |  | 7 | 0 | 7.2992 | 0.86327 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 227.7472 | 499972.6391 | 5.6989 | 5.7028 | 0.06746% | 504421.4537 |  | 0 |  |  | 0.86327 | 7.2992 |  | 8 | 0 | 7.2992 | 0.86327 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.1047% |  |  | 0 |  |  | 0.86327 | 7.2992 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.02125 |  | 0 |  |  | 0.86327 | 7.2992 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 0.86327 |  | 0 |  |  | 0.86327 | 7.2992 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99991 |  |  |  |  | 0.86327 | 7.2992 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 0.86327 | 7.2992 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MECSB/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MECSB | MECSB/CH#
[9 empty] | #VALUE! | 0 | 1 | 7.2992 | 7.2992
[9 empty] | #VALUE! | 0 | 1 | 7.2992 | 7.2992
[9 empty] | #VALUE! | 0 | 1 | 7.2992 | 7.2992
[9 empty] | #VALUE! | 0 | 1 | 7.2992 | 7.2992
[9 empty] | #VALUE! | 0 | 1 | 7.2992 | 7.2992
#VALUE!
[9 empty] | #VALUE! | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  |  | #VALUE! | Determination of New MECSB
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MECSB
[8 empty] | 105 MEAN CH# |  |  | 1 | 7.2992
[8 empty] | to 256 REL. CH# |  |  | 1 | 7.2992
[8 empty] | CONVERSION |  |  | 1 | 7.2992
[8 empty] | 105 CH# | 256 CH# |  | 1 | 7.2992
[8 empty] | 282 | 115.7625 |  | 1 | 7.2992
[8 empty] | 561 | 129.8753 |  | 1 | 7.2992
[8 empty] | 881 | 139.136 |  | 1 | 7.2992
[8 empty] | 1672 | 152.2824 |  | 1 | 7.2992
[8 empty] | 4816 | 173.9893
[8 empty] | 12294 | 193.2184
[8 empty] | 34838 | 214.5902
[8 empty] | 66151 | 227.7472
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#

## MEBFP
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MEBFP | MEBFP LOG | CALC. | RESIDUAL | CALC. MEBFP | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MEBFP | MEBFP LOG | CALC. | RESIDUAL | CALC. MEBFP | CROSS CALIBRATION TABLE
| 1 | 110.6682 |  |  | 2.9877 |  | 972.1153 | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MEBFP |  | 1 | 0 | 3.7673 | 0.57603 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 128.9016 | 2584.069 | 3.4123 | 3.3851 | 0.8048% | 2426.9541 | CONVERSION |  |  |  | 0.57603 | 3.7673 |  | 2 | 0 | 3.7673 | 0.57603 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MEBFP
| 3 | 143.4095 | 5291.2727 | 3.7236 | 3.7012 | 0.6036% | 5025.9519 | 1024 CH# | 256 CH# |  |  | 0.57603 | 3.7673 |  | 3 | 0 | 3.7673 | 0.57603 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 4 | 160.2221 | 11909.3206 | 4.0759 | 4.0676 | 0.2037% | 11684.2279 |  | 0 |  |  | 0.57603 | 3.7673 |  | 4 | 0 | 3.7673 | 0.57603 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 184.3046 | 38804.9158 | 4.5889 | 4.5924 | 0.07667% | 39120.8172 |  | 0 |  |  | 0.57603 | 3.7673 |  | 5 | 0 | 3.7673 | 0.57603 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 204.9402 | 107470.7906 | 5.0313 | 5.0421 | 0.2144% | 110179.4508 |  | 0 |  |  | 0.57603 | 3.7673 |  | 6 | 0 | 3.7673 | 0.57603 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 228.7537 | 359515.7482 | 5.5557 | 5.561 | 0.09586% | 363955.7173 |  | 0 |  |  | 0.57603 | 3.7673 |  | 7 | 0 | 3.7673 | 0.57603 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 242.2903 | 736906.7773 | 5.8674 | 5.856 | 0.1942% | 717856.3695 |  | 0 |  |  | 0.57603 | 3.7673 |  | 8 | 0 | 3.7673 | 0.57603 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.157% |  |  | 0 |  |  | 0.57603 | 3.7673 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.021792 |  | 0 |  |  | 0.57603 | 3.7673 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 0.57603 |  | 0 |  |  | 0.57603 | 3.7673 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99983 |  |  |  |  | 0.57603 | 3.7673 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 0.57603 | 3.7673 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MEBFP/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MEBFP | MEBFP/CH#
[9 empty] | #VALUE! | 0 | 1 | 3.7673 | 3.7673
[9 empty] | #VALUE! | 0 | 1 | 3.7673 | 3.7673
[9 empty] | #VALUE! | 0 | 1 | 3.7673 | 3.7673
[9 empty] | #VALUE! | 0 | 1 | 3.7673 | 3.7673
[9 empty] | #VALUE! | 0 | 1 | 3.7673 | 3.7673
#VALUE!
[9 empty] | #VALUE! | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  |  | #VALUE! | Determination of New MEBFP
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MEBFP
[8 empty] | 105 MEAN CH# |  |  | 1 | 3.7673
[8 empty] | to 256 REL. CH# |  |  | 1 | 3.7673
[8 empty] | CONVERSION |  |  | 1 | 3.7673
[8 empty] | 105 CH# | 256 CH# |  | 1 | 3.7673
[8 empty] | 220 | 110.6682 |  | 1 | 3.7673
[8 empty] | 535 | 128.9016 |  | 1 | 3.7673
[8 empty] | 1085 | 143.4095 |  | 1 | 3.7673
[8 empty] | 2462 | 160.2221 |  | 1 | 3.7673
[8 empty] | 7962 | 184.3046
[8 empty] | 21767 | 204.9402
[8 empty] | 69477 | 228.7537
[8 empty] | 134387 | 242.2903
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#

## MEFL
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MEFL | MEFL LOG | CALC. | RESIDUAL | CALC. MEFL | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MEFL | MEFL LOG | CALC. | RESIDUAL | CALC. MEFL | CROSS CALIBRATION TABLE
| 1 | 49.5691 |  |  | 2.3802 |  | 239.9718 | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MEFL |  | 1 | 0 | 38.8695 | 1.5896 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 83.6386 | 806 | 2.9063 | 2.9235 | 0.5877% | 838.5287 | CONVERSION |  |  |  | 1.5896 | 38.8695 |  | 2 | 0 | 38.8695 | 1.5896 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MEFL
| 3 | 108.8229 | 2159 | 3.3343 | 3.3252 | 0.2732% | 2114.3088 | 1024 CH# | 256 CH# |  |  | 1.5896 | 38.8695 |  | 3 | 0 | 38.8695 | 1.5896 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 4 | 134.8614 | 5640 | 3.7513 | 3.7404 | 0.2897% | 5501.0285 |  | 0 |  |  | 1.5896 | 38.8695 |  | 4 | 0 | 38.8695 | 1.5896 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 169.1942 | 19900 | 4.2989 | 4.288 | 0.2531% | 19408.8556 |  | 0 |  |  | 1.5896 | 38.8695 |  | 5 | 0 | 38.8695 | 1.5896 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 196.8982 | 52630 | 4.7212 | 4.7298 | 0.1819% | 53682.9874 |  | 0 |  |  | 1.5896 | 38.8695 |  | 6 | 0 | 38.8695 | 1.5896 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 228.8978 | 172155 | 5.2359 | 5.2402 | 0.08132% | 173852.555 |  | 0 |  |  | 1.5896 | 38.8695 |  | 7 | 0 | 38.8695 | 1.5896 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 247.6743 | 345870 | 5.5389 | 5.5396 | 0.01309% | 346447.8256 |  | 0 |  |  | 1.5896 | 38.8695 |  | 8 | 0 | 38.8695 | 1.5896 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.24% |  |  | 0 |  |  | 1.5896 | 38.8695 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.015948 |  | 0 |  |  | 1.5896 | 38.8695 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 1.5896 |  | 0 |  |  | 1.5896 | 38.8695 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99988 |  |  |  |  | 1.5896 | 38.8695 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 1.5896 | 38.8695 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MEFL/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MEFL | MEFL/CH#
[8 empty] | 5.95 | 49.5691 | 0 | 1 | 38.8695 | 38.8695
[8 empty] | 20.27 | 83.6386 | 0 | 1 | 38.8695 | 38.8695
[8 empty] | 50.16 | 108.8229 | 0 | 1 | 38.8695 | 38.8695
[8 empty] | 128 | 134.8614 | 0 | 1 | 38.8695 | 38.8695
[8 empty] | 440.21 | 169.1942 | 0 | 1 | 38.8695 | 38.8695
[8 empty] | 1192.71 | 196.8982
[8 empty] | 3771.62 | 228.8978 | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  | 7411.58 | 247.6743 | Determination of New MEFL
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MEFL
[8 empty] | 105 MEAN CH# |  |  | 1 | 38.8695
[8 empty] | to 256 REL. CH# |  |  | 1 | 38.8695
[8 empty] | CONVERSION |  |  | 1 | 38.8695
[8 empty] | 105 CH# | 256 CH# |  | 1 | 38.8695
[8 empty] | 25 | 66.046 |  | 1 | 38.8695
[8 empty] | 152 | 103.0816 |  | 1 | 38.8695
[8 empty] | 414 | 123.6407 |  | 1 | 38.8695
[8 empty] | 1078 | 143.2767 |  | 1 | 38.8695
[8 empty] | 3868 | 169.4915
[8 empty] | 10282 | 189.5514
[8 empty] | 32712 | 213.2982
[8 empty] | 65004 | 227.3883
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#

## MEPE
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MEPE | MEPE LOG | CALC. | RESIDUAL | CALC. MEPE | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  | PEAK # | CH # | MEPE | MEPE LOG | CALC. | RESIDUAL | CALC. MEPE | CROSS CALIBRATION TABLE
| 1 | 31.4471 |  |  | 1.9855 |  | 96.7264 | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MEPE | 1 | 0 | 30.4924 | 1.4842 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 71.6112 | 409 | 2.6117 | 2.6259 | 0.5387% | 422.542 | CONVERSION |  |  |  | 1.4842 | 30.4924 | 2 | 0 | 30.4924 | 1.4842 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MEPE
| 3 | 100.7528 | 1250 | 3.0969 | 3.0905 | 0.2085% | 1231.5889 | 1024 CH# | 256 CH# |  |  | 1.4842 | 30.4924 | 3 | 0 | 30.4924 | 1.4842 | #VALUE! | #VALUE! | #VALUE! |  | 200 | #VALUE! | #VALUE!
| 4 | 128 | 3428 | 3.535 | 3.5249 | 0.2888% | 3348.5735 |  | 0 |  |  | 1.4842 | 30.4924 | 4 | 0 | 30.4924 | 1.4842 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 163.0087 | 12229 | 4.0874 | 4.083 | 0.1077% | 12105.8124 |  | 0 |  |  | 1.4842 | 30.4924 | 5 | 0 | 30.4924 | 1.4842 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 191.3717 | 34294 | 4.5352 | 4.5352 | 0.0008962% | 34290.7908 |  | 0 |  |  | 1.4842 | 30.4924 | 6 | 0 | 30.4924 | 1.4842 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 223.999 | 113118 | 5.0535 | 5.0553 | 0.03585% | 113591.0604 |  | 0 |  |  | 1.4842 | 30.4924 | 7 | 0 | 30.4924 | 1.4842 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 246.4686 | 256134 | 5.4085 | 5.4136 | 0.09428% | 259161.8285 |  | 0 |  |  | 1.4842 | 30.4924 | 8 | 0 | 30.4924 | 1.4842 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.1821% |  |  | 0 |  |  | 1.4842 | 30.4924 |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.015943 |  | 0 |  |  | 1.4842 | 30.4924 | [5 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 1.4842 |  | 0 |  |  | 1.4842 | 30.4924 | [5 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99994 |  |  |  |  | 1.4842 | 30.4924 | [5 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 1.4842 | 30.4924 | [9 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [14 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[7 empty] | 21 | to 256 REL. CH# |  | CAN BE COPIED TO
[7 empty] | 56.5 | CONVERSION |  | THE CH# COLUMN BY | [9 empty] | 21
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS | [9 empty] | 56.5
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MEPE/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MEPE | MEPE/CH#
[8 empty] | 3.1 | 31.4471 | 0 | 1 | 30.4924 | 30.4924
[8 empty] | 13.15 | 71.6112 | 0 | 1 | 30.4924 | 30.4924
[8 empty] | 37.52 | 100.7528 | 0 | 1 | 30.4924 | 30.4924
[8 empty] | 100 | 128 | 0 | 1 | 30.4924 | 30.4924
[8 empty] | 352.38 | 163.0087 | 0 | 1 | 30.4924 | 30.4924
[8 empty] | 977.65 | 191.3717
[8 empty] | 3162.16 | 223.999 | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  | 7096.95 | 246.4686 | Determination of New MEPE
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MEPE
[8 empty] | 105 MEAN CH# |  |  | 1 | 30.4924
[8 empty] | to 256 REL. CH# |  |  | 1 | 30.4924
[8 empty] | CONVERSION |  |  | 1 | 30.4924
[8 empty] | 105 CH# | 256 CH# |  | 1 | 30.4924
[8 empty] | 81 | 90.1668 |  | 1 | 30.4924
[8 empty] | 176 | 106.0897 |  | 1 | 30.4924
[8 empty] | 419 | 123.887 |  | 1 | 30.4924
[8 empty] | 1031 | 142.362 |  | 1 | 30.4924
[8 empty] | 3439 | 167.0794
[8 empty] | 9604 | 188.1517
[8 empty] | 31498 | 212.5223
[8 empty] | 74628 | 230.2212
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#
Note on T5: RCP-30-5 Relative Channel Numbers obtained with your Instrument in 256 Channel Scale.

## MEPTR
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
[8 empty] | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MEPTR | MEPTR LOG | CALC. | RESIDUAL | CALC. MEPTR | CROSS CALIBRATION TABLE
| PEAK # | CH # | MEPTR | MEPTR LOG | CALC. | RESIDUAL | CALC. MEPTR | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MEPTR |  | 1 | 0 | 16.6027 | 1.2202 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 1 | 19.8163 |  |  | 1.5461 |  | 35.166 | CONVERSION |  |  |  | 1.2202 | 16.6027 |  | 2 | 0 | 16.6027 | 1.2202 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MEPTR
| 2 | 62.2506 | 171 | 2.233 | 2.2441 | 0.4944% | 175.4251 | 1024 CH# | 256 CH# |  |  | 1.2202 | 16.6027 |  | 3 | 0 | 16.6027 | 1.2202 | #VALUE! | #VALUE! | #VALUE! |  | 200 | #VALUE! | #VALUE!
| 3 | 91.669 | 538 | 2.7308 | 2.728 | 0.103% | 534.5306 |  | 0 |  |  | 1.2202 | 16.6027 |  | 4 | 0 | 16.6027 | 1.2202 | #VALUE! | #VALUE! | #VALUE! |  | 125 | #VALUE! | #VALUE!
| 4 | 118.5665 | 1514 | 3.1801 | 3.1704 | 0.3071% | 1480.4376 |  | 0 |  |  | 1.2202 | 16.6027 |  | 5 | 0 | 16.6027 | 1.2202 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 153.8601 | 5659 | 3.7527 | 3.7509 | 0.04886% | 5635.1708 |  | 0 |  |  | 1.2202 | 16.6027 |  | 6 | 0 | 16.6027 | 1.2202 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 182.7838 | 16972 | 4.2297 | 4.2267 | 0.07289% | 16852.0329 |  | 0 |  |  | 1.2202 | 16.6027 |  | 7 | 0 | 16.6027 | 1.2202 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 218.2715 | 64615 | 4.8103 | 4.8104 | 0.0005987% | 64619.2849 |  | 0 |  |  | 1.2202 | 16.6027 |  | 8 | 0 | 16.6027 | 1.2202 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 246.2712 | 183897 | 5.2646 | 5.2709 | 0.1202% | 186598.8778 |  | 0 |  |  | 1.2202 | 16.6027 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.1639% |  |  | 0 |  |  | 1.2202 | 16.6027 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.016448 |  | 0 |  |  | 1.2202 | 16.6027 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 1.2202 |  |  |  |  | 1.2202 | 16.6027 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99996 |  |  |  |  | 1.2202 | 16.6027 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MEPTR/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MEPTR | MEPTR/CH#
[8 empty] | 2.04 | 19.8163 | 0 | 1 | 16.6027 | 16.6027
[8 empty] | 9.39 | 62.2506 | 0 | 1 | 16.6027 | 16.6027
[8 empty] | 27.06 | 91.669 | 0 | 1 | 16.6027 | 16.6027
[8 empty] | 71.22 | 118.5665 | 0 | 1 | 16.6027 | 16.6027
[8 empty] | 253.55 | 153.8601 | 0 | 1 | 16.6027 | 16.6027
[8 empty] | 717.79 | 182.7838
[8 empty] | 2573.31 | 218.2715 | TABLE NO. 6
[8 empty] | 7046.74 | 246.2712 | Determination of New MEPTR
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  |  |  | values for RCP-30-5A
Acceptable:
[8 empty] | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MEPTR
Action taken if not linear: |  |  | By: |  |  | Date: |  | 105 MEAN CH# |  |  | 1 | 16.6027
[8 empty] | to 256 REL. CH# |  |  | 1 | 16.6027
[8 empty] | CONVERSION |  |  | 1 | 16.6027
[8 empty] | 105 CH# | 256 CH# |  | 1 | 16.6027
[8 empty] | 2 | 14.2222 |  | 1 | 16.6027
[8 empty] | 57 | 82.9567 |  | 1 | 16.6027
[8 empty] | 201 | 108.815 |  | 1 | 16.6027
[8 empty] | 563 | 129.9483 |  | 1 | 16.6027
[8 empty] | 1970 | 155.6477
[8 empty] | 5687 | 177.4002
[8 empty] | 19483 | 202.6656
[8 empty] | 49482 | 221.7901
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#

## MECY
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MECY | MECY LOG | CALC. | RESIDUAL | CALC. MECY | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MECY | MECY LOG | CALC. | RESIDUAL | CALC. MECY | CROSS CALIBRATION TABLE
| 1 | 25.7997 |  |  | 2.5754 |  | 376.2009 | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MECY |  | 1 | 0 | 149.64 | 2.175 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 60.7297 | 1276 | 3.1059 | 3.1175 | 0.373% | 1310.6299 | CONVERSION |  |  |  | 2.175 | 149.64 |  | 2 | 0 | 149.64 | 2.175 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MECY
| 3 | 88.2407 | 3603 | 3.5567 | 3.5444 | 0.3458% | 3502.7477 | 1024 CH# | 256 CH# |  |  | 2.175 | 149.64 |  | 3 | 0 | 149.64 | 2.175 | #VALUE! | #VALUE! | #VALUE! |  | 200 | #VALUE! | #VALUE!
| 4 | 115.4431 | 9520 | 3.9786 | 3.9665 | 0.3048% | 9258.6642 |  | 0 |  |  | 2.175 | 149.64 |  | 4 | 0 | 149.64 | 2.175 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 153.0074 | 35518 | 4.5504 | 4.5495 | 0.02109% | 35439.6335 |  | 0 |  |  | 2.175 | 149.64 |  | 5 | 0 | 149.64 | 2.175 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 183.6822 | 102755 | 5.0118 | 5.0255 | 0.2729% | 106051.2275 |  | 0 |  |  | 2.175 | 149.64 |  | 6 | 0 | 149.64 | 2.175 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 216.5642 | 328185 | 5.5161 | 5.5358 | 0.3554% | 343394.4564 |  | 0 |  |  | 2.175 | 149.64 |  | 7 | 0 | 149.64 | 2.175 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 246.6389 | 1052496 | 6.0222 | 6.0025 | 0.3284% | 1005789.4104 |  | 0 |  |  | 2.175 | 149.64 |  | 8 | 0 | 149.64 | 2.175 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.2859% |  |  | 0 |  |  | 2.175 | 149.64 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.015518 |  | 0 |  |  | 2.175 | 149.64 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 2.175 |  | 0 |  |  | 2.175 | 149.64 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99979 |  |  |  |  | 2.175 | 149.64 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 2.175 | 149.64 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MECY/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MECY | MECY/CH#
[8 empty] | 2.53 | 25.7997 | 0 | 1 | 149.64 | 149.64
[8 empty] | 8.89 | 60.7297 | 0 | 1 | 149.64 | 149.64
[8 empty] | 23.92 | 88.2407 | 0 | 1 | 149.64 | 149.64
[8 empty] | 63.65 | 115.4431 | 0 | 1 | 149.64 | 149.64
[8 empty] | 245.89 | 153.0074 | 0 | 1 | 149.64 | 149.64
[8 empty] | 741.37 | 183.6822
[8 empty] | 2420 | 216.5642 | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  | 7140.57 | 246.6389 | Determination of New MECY
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MECY
[8 empty] | 105 MEAN CH# |  |  | 1 | 149.64
[8 empty] | to 256 REL. CH# |  |  | 1 | 149.64
[8 empty] | CONVERSION |  |  | 1 | 149.64
[8 empty] | 105 CH# | 256 CH# |  | 1 | 149.64
[8 empty] | 7 | 39.9268 |  | 1 | 149.64
[8 empty] | 84 | 90.913 |  | 1 | 149.64
[8 empty] | 297 | 116.8259 |  | 1 | 149.64
[8 empty] | 875 | 138.9958 |  | 1 | 149.64
[8 empty] | 3234 | 165.8184
[8 empty] | 8665 | 186.0407
[8 empty] | 21089 | 204.2909
[8 empty] | 55019 | 223.9665
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#

## MEPCY7
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MEPCY7 | MEPCY7 LOG | CALC. | RESIDUAL | CALC. MEPCY7 | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MEPCY7 | MEPCY7 LOG | CALC. | RESIDUAL | CALC. MEPCY7 | CROSS CALIBRATION TABLE
| 1 | 50.488 | [5 empty] | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MEPCY7 |  | 1 | 0 | 567.5846 | 2.754 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 64.8486 | [5 empty] | CONVERSION |  |  |  | 2.754 | 567.5846 |  | 2 | 0 | 567.5846 | 2.754 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MEPCY7
| 3 | 80.6446 | [5 empty] | 1024 CH# | 256 CH# |  |  | 2.754 | 567.5846 |  | 3 | 0 | 567.5846 | 2.754 | #VALUE! | #VALUE! | #VALUE! |  | 200 | #VALUE! | #VALUE!
| 4 | 102.5874 | 9726.6114 | 3.988 | 3.988 | 0% | 9726.6114 |  | 0 |  |  | 2.754 | 567.5846 |  | 4 | 0 | 567.5846 | 2.754 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 136.922 | 25173.42 | 4.4009 | 4.4009 | 0% | 25173.42 |  | 0 |  |  | 2.754 | 567.5846 |  | 5 | 0 | 567.5846 | 2.754 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 166.4845 | 57085.1374 | 4.7565 | 4.7565 | 0% | 57085.1374 |  | 0 |  |  | 2.754 | 567.5846 |  | 6 | 0 | 567.5846 | 2.754 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 198.022 | 136728.333 | 5.1359 | 5.1359 | 0% | 136728.333 |  | 0 |  |  | 2.754 | 567.5846 |  | 7 | 0 | 567.5846 | 2.754 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 229.2489 | 324682.7028 | 5.5115 | 5.5115 | 0% | 324682.7028 |  | 0 |  |  | 2.754 | 567.5846 |  | 8 | 0 | 567.5846 | 2.754 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0% |  |  | 0 |  |  | 2.754 | 567.5846 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.012028 |  | 0 |  |  | 2.754 | 567.5846 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 2.754 |  | 0 |  |  | 2.754 | 567.5846 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 1 |  |  |  |  | 2.754 | 567.5846 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 2.754 | 567.5846 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MEAP/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MEPCY7 | MEPCY7/CH#
[8 empty] | 6.15 | 50.488 | 0 | 1 | 567.5846 | 567.5846
[8 empty] | 10.31 | 64.8486 | 0 | 1 | 567.5846 | 567.5846
[8 empty] | 18.2 | 80.6446 | 0 | 1 | 567.5846 | 567.5846
[8 empty] | 40.08 | 102.5874 | 0 | 1 | 567.5846 | 567.5846
[8 empty] | 137.85 | 136.922 | 0 | 1 | 567.5846 | 567.5846
[8 empty] | 399.32 | 166.4845
[8 empty] | 1241.92 | 198.022 | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  | 3819.57 | 229.2489 | Determination of New MEPCY7
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MEPCY7
[8 empty] | 105 MEAN CH# |  |  | 1 | 567.5846
[8 empty] | to 256 REL. CH# |  |  | 1 | 567.5846
[8 empty] | CONVERSION |  |  | 1 | 567.5846
[8 empty] | 105 CH# | 256 CH# |  | 1 | 567.5846
[8 empty] | 8 | 42.6667 |  | 1 | 567.5846
[8 empty] | 75 | 88.5876 |  | 1 | 567.5846
[8 empty] | 245 | 112.8766 |  | 1 | 567.5846
[8 empty] | 745 | 135.6956 |  | 1 | 567.5846
[8 empty] | 2913 | 163.6735
[8 empty] | 7831 | 183.9642
[8 empty] | 16545 | 199.3118
[8 empty] | 43215 | 219.0115
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#

## MEAP
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MEAP | MEAP LOG | CALC. | RESIDUAL | CALC. MEAP | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MEAP | MEAP LOG | CALC. | RESIDUAL | CALC. MEAP | CROSS CALIBRATION TABLE
| 1 | 51.7255 |  |  | 2.3477 |  | 222.6921 | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MEAP |  | 1 | 0 | 34.5693 | 1.5387 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 101.5704 | 1366 | 3.1355 | 3.1273 | 0.2605% | 1340.6178 | CONVERSION |  |  |  | 1.5387 | 34.5693 |  | 2 | 0 | 34.5693 | 1.5387 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MEAP
| 3 | 136.8413 | 4706 | 3.6727 | 3.679 | 0.1715% | 4774.8564 | 1024 CH# | 256 CH# |  |  | 1.5387 | 34.5693 |  | 3 | 0 | 34.5693 | 1.5387 | #VALUE! | #VALUE! | #VALUE! |  | 200 | #VALUE! | #VALUE!
| 4 | 167.8217 | 14822 | 4.1709 | 4.1635 | 0.1776% | 14571.7147 |  | 0 |  |  | 1.5387 | 34.5693 |  | 4 | 0 | 34.5693 | 1.5387 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 206.69 | 58161 | 4.7646 | 4.7714 | 0.1425% | 59078.7115 |  | 0 |  |  | 1.5387 | 34.5693 |  | 5 | 0 | 34.5693 | 1.5387 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 231.9806 | 140920 | 5.149 | 5.167 | 0.3487% | 146889.4257 |  | 0 |  |  | 1.5387 | 34.5693 |  | 6 | 0 | 34.5693 | 1.5387 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 241.2993 | 204810 | 5.3114 | 5.3127 | 0.02613% | 205465.7388 |  | 0 |  |  | 1.5387 | 34.5693 |  | 7 | 0 | 34.5693 | 1.5387 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 248.755 | 280646 | 5.4482 | 5.4293 | 0.3464% | 268750.7596 |  | 0 |  |  | 1.5387 | 34.5693 |  | 8 | 0 | 34.5693 | 1.5387 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.2083% |  |  | 0 |  |  | 1.5387 | 34.5693 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.015641 |  | 0 |  |  | 1.5387 | 34.5693 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 1.5387 |  | 0 |  |  | 1.5387 | 34.5693 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99926 |  |  |  |  | 1.5387 | 34.5693 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 1.5387 | 34.5693 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MEAP/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MEAP | MEAP/CH#
[8 empty] | 6.43 | 51.7255 | 0 | 1 | 34.5693 | 34.5693
[8 empty] | 38.64 | 101.5704 | 0 | 1 | 34.5693 | 34.5693
[8 empty] | 137.45 | 136.8413 | 0 | 1 | 34.5693 | 34.5693
[8 empty] | 419 | 167.8217 | 0 | 1 | 34.5693 | 34.5693
[8 empty] | 1696.41 | 206.69 | 0 | 1 | 34.5693 | 34.5693
[8 empty] | 4214.03 | 231.9806
[8 empty] | 5892.53 | 241.2993 | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  | 7705.42 | 248.755 | Determination of New MEAP
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MEAP
[8 empty] | 105 MEAN CH# |  |  | 1 | 34.5693
[8 empty] | to 256 REL. CH# |  |  | 1 | 34.5693
[8 empty] | CONVERSION |  |  | 1 | 34.5693
[8 empty] | 105 CH# | 256 CH# |  | 1 | 34.5693
[8 empty] | 35 | 72.9498 |  | 1 | 34.5693
[8 empty] | 227 | 111.3109 |  | 1 | 34.5693
[8 empty] | 773 | 136.4526 |  | 1 | 34.5693
[8 empty] | 2401 | 159.7073 |  | 1 | 34.5693
[8 empty] | 9283 | 187.4542
[8 empty] | 22295 | 205.4319
[8 empty] | 31325 | 212.4093
[8 empty] | 42061 | 218.4561
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#

## MEAXL700
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MEALX700 | MEAXL700 LOG | CALC. | RESIDUAL | CALC. MEAXL700 | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MEALX700 | MEAXL700 LOG | CALC. | RESIDUAL | CALC. MEAXL700 | CROSS CALIBRATION TABLE
| 1 | 52.6285 |  |  | 2.2555 |  | 180.0765 | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MEAXL700 |  | 1 | 0 | 14.5272 | 1.1622 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 91.633 | 1180.8957 | 3.0722 | 3.0657 | 0.2119% | 1163.3656 | CONVERSION |  |  |  | 1.1622 | 14.5272 |  | 2 | 0 | 14.5272 | 1.1622 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MEAXL700
| 3 | 116.8259 | 3930.9658 | 3.5945 | 3.5891 | 0.1515% | 3882.0433 | 1024 CH# | 256 CH# |  |  | 1.1622 | 14.5272 |  | 3 | 0 | 14.5272 | 1.1622 | #VALUE! | #VALUE! | #VALUE! |  | 200 | #VALUE! | #VALUE!
| 4 | 140.1581 | 11270.61 | 4.0519 | 4.0738 | 0.5353% | 11850.9321 |  | 0 |  |  | 1.1622 | 14.5272 |  | 4 | 0 | 14.5272 | 1.1622 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 168.0751 | 45502.2218 | 4.658 | 4.6537 | 0.0934% | 45049.0803 |  | 0 |  |  | 1.1622 | 14.5272 |  | 5 | 0 | 14.5272 | 1.1622 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 186.2855 | 109122.0814 | 5.0379 | 5.032 | 0.1179% | 107641.0044 |  | 0 |  |  | 1.1622 | 14.5272 |  | 6 | 0 | 14.5272 | 1.1622 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 193.8753 | 151781.0757 | 5.1812 | 5.1896 | 0.1624% | 154755.3539 |  | 0 |  |  | 1.1622 | 14.5272 |  | 7 | 0 | 14.5272 | 1.1622 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 202.4869 | 237988.0139 | 5.3766 | 5.3685 | 0.1493% | 233634.9027 |  | 0 |  |  | 1.1622 | 14.5272 |  | 8 | 0 | 14.5272 | 1.1622 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.2031% |  |  | 0 |  |  | 1.1622 | 14.5272 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.020773 |  | 0 |  |  | 1.1622 | 14.5272 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 1.1622 |  | 0 |  |  | 1.1622 | 14.5272 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99984 |  |  |  |  | 1.1622 | 14.5272 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 1.1622 | 14.5272 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MEAXL700/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MEAXL700 | MEAXL700/CH#
[9 empty] | #VALUE! | 0 | 1 | 14.5272 | 14.5272
[9 empty] | #VALUE! | 0 | 1 | 14.5272 | 14.5272
[9 empty] | #VALUE! | 0 | 1 | 14.5272 | 14.5272
[9 empty] | #VALUE! | 0 | 1 | 14.5272 | 14.5272
[9 empty] | #VALUE! | 0 | 1 | 14.5272 | 14.5272
#VALUE!
[9 empty] | #VALUE! | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  |  | #VALUE! | Determination of New MEAXL700
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MEAXL700
[8 empty] | 105 MEAN CH# |  |  | 1 | 14.5272
[8 empty] | to 256 REL. CH# |  |  | 1 | 14.5272
[8 empty] | CONVERSION |  |  | 1 | 14.5272
[8 empty] | 105 CH# | 256 CH# |  | 1 | 14.5272
[8 empty] | 13 | 52.6285 |  | 1 | 14.5272
[8 empty] | 87 | 91.633 |  | 1 | 14.5272
[8 empty] | 297 | 116.8259 |  | 1 | 14.5272
[8 empty] | 926 | 140.1581 |  | 1 | 14.5272
[8 empty] | 3610 | 168.0751
[8 empty] | 8769 | 186.2855
[8 empty] | 12694 | 193.8753
[8 empty] | 19314 | 202.4869
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#

## MEAPCY7
TYPE THE NAME OF YOUR LAB HERE
| PMT LINEARITY QC RECORD | [12 empty] | Normalization Graph for Instrument B
[8 empty] | TABLE NO. 1 |  | CROSS CALIBRATION TABLE
| PEAK # | CH # | MEAPCY7 | MEAPCY7 LOG | CALC. | RESIDUAL | CALC. MEAPCY7 | 1024 REL. CH# |  | FOR UNKNOWN SAMPLES |  |  |  |  | PEAK # | CH # | MEAPCY7 | MEAPCY7 LOG | CALC. | RESIDUAL | CALC. MEAPCY7 | CROSS CALIBRATION TABLE
| 1 | 53.887 |  |  | 3.6686 |  | 4661.8736 | to 256 REL. CH# |  | Sample | CH# | Calc. | Calc. MEAPCY7 |  | 1 | 0 | 1384.5013 | 3.1413 | #VALUE! | #VALUE! | #VALUE! | FOR NORMALIZATION GRAPH
| 2 | 98.693 | 12540.2912 | 4.0983 | 4.107 | 0.211% | 12793.0013 | CONVERSION |  |  |  | 3.1413 | 1384.5013 |  | 2 | 0 | 1384.5013 | 3.1413 | #VALUE! | #VALUE! | #VALUE! | Sample | CH# | Calc. | Calc. MEAPCY7
| 3 | 132.7929 | 27790.4244 | 4.4439 | 4.4406 | 0.07355% | 27582.2202 | 1024 CH# | 256 CH# |  |  | 3.1413 | 1384.5013 |  | 3 | 0 | 1384.5013 | 3.1413 | #VALUE! | #VALUE! | #VALUE! |  | 200 | #VALUE! | #VALUE!
| 4 | 163.7035 | 57114.753 | 4.7567 | 4.7431 | 0.2882% | 55345.1564 |  | 0 |  |  | 3.1413 | 1384.5013 |  | 4 | 0 | 1384.5013 | 3.1413 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 5 | 204.0765 | 137792.7954 | 5.1392 | 5.1381 | 0.02162% | 137440.8227 |  | 0 |  |  | 3.1413 | 1384.5013 |  | 5 | 0 | 1384.5013 | 3.1413 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 6 | 230.8701 | 245420.4198 | 5.3899 | 5.4003 | 0.1921% | 251351.8609 |  | 0 |  |  | 3.1413 | 1384.5013 |  | 6 | 0 | 1384.5013 | 3.1413 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 7 | 242.7222 | 312870.989 | 5.4954 | 5.5163 | 0.3786% | 328285.4669 |  | 0 |  |  | 3.1413 | 1384.5013 |  | 7 | 0 | 1384.5013 | 3.1413 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
| 8 | 247.8015 | 387104.3884 | 5.5878 | 5.566 | 0.3931% | 368087.3621 |  | 0 |  |  | 3.1413 | 1384.5013 |  | 8 | 0 | 1384.5013 | 3.1413 | #VALUE! | #VALUE! | #VALUE! |  |  | #VALUE! | #VALUE!
|  |  |  | Ave Residual |  | 0.2226% |  |  | 0 |  |  | 3.1413 | 1384.5013 |  |  |  |  | Ave Residual |  | #VALUE! |  |  |  | #VALUE! | #VALUE!
[6 empty] | Slope: | 0.0097847 |  | 0 |  |  | 3.1413 | 1384.5013 | [6 empty] | Slope: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Intercept: | 3.1413 |  | 0 |  |  | 3.1413 | 1384.5013 | [6 empty] | Intercept: | #VALUE! |  |  | #VALUE! | #VALUE!
[6 empty] | Rsq: | 0.99933 |  |  |  |  | 3.1413 | 1384.5013 | [6 empty] | Rsq: | #VALUE! |  |  | #VALUE! | #VALUE!
[12 empty] | 3.1413 | 1384.5013 | [10 empty] | #VALUE! | #VALUE!
[8 empty] | TABLE NO. 2 | [15 empty] | #VALUE! | #VALUE!
[8 empty] | 1024 MEAN CH# |  | CONVERTED VALUES
[8 empty] | to 256 REL. CH# |  | CAN BE COPIED TO
[8 empty] | CONVERSION |  | THE CH# COLUMN BY
[8 empty] | 1024 CH# | 256 CH# | SELECTING THE CELLS
[9 empty] | #VALUE! | TO BE COPIED; THEN,
[9 empty] | #VALUE! | CLICKING [COPY],
[9 empty] | #VALUE! | [PASTE SPECIAL], AND
[9 empty] | #VALUE! | FINALLY [PASTE VALUE]
#VALUE!
#VALUE!
#VALUE!
#VALUE!
[8 empty] | TABLE NO. 3 |  | TABLE NO. 5
[8 empty] | 104 MEAN CH# |  | Determination of the
[8 empty] | to 256 REL. CH# |  | MEAPCY7/CH# for the
[8 empty] | CONVERSION |  | normalization cell or particle
[8 empty] | 104 CH# | 256 CH# | 256 CH# | 104 CH# | Calc. MEAPC7 | MEAPCY7/CH#
[8 empty] | 6.95 | 53.887 | 0 | 1 | 1384.5013 | 1384.5013
[8 empty] | 34.84 | 98.693 | 0 | 1 | 1384.5013 | 1384.5013
[8 empty] | 118.82 | 132.7929 | 0 | 1 | 1384.5013 | 1384.5013
[8 empty] | 361.3 | 163.7035 | 0 | 1 | 1384.5013 | 1384.5013
[8 empty] | 1544.17 | 204.0765 | 0 | 1 | 1384.5013 | 1384.5013
[8 empty] | 4048.97 | 230.8701
[8 empty] | 6202.04 | 242.7222 | TABLE NO. 6
Rainbow Calibration Particles (RCP-30-5A) |  |  |  | Lot No.: |  | File # |  | 7445.58 | 247.8015 | Determination of New MEAPCY7
Acceptable: | [9 empty] | values for RCP-30-5A
Action taken if not linear: |  |  | By: |  |  | Date: |  | TABLE NO. 4 |  | 256 CH# | 104 CH# | New MEAPCY7
[8 empty] | 105 MEAN CH# |  |  | 1 | 1384.5013
[8 empty] | to 256 REL. CH# |  |  | 1 | 1384.5013
[8 empty] | CONVERSION |  |  | 1 | 1384.5013
[8 empty] | 105 CH# | 256 CH# |  | 1 | 1384.5013
[8 empty] | 30 | 69.7869 |  | 1 | 1384.5013
[8 empty] | 232 | 111.758 |  | 1 | 1384.5013
[8 empty] | 812 | 137.4626 |  | 1 | 1384.5013
[8 empty] | 2524 | 160.7324 |  | 1 | 1384.5013
[8 empty] | 10098 | 189.1809
[8 empty] | 25055 | 207.8266
[8 empty] | 36721 | 215.6703
[8 empty] | 51343 | 222.5476
Note on J20: Use this Table to convert Beckman-Coulter 1024 Mean CH # to 256 Relative CH#
Note on J35: Use this Table to convert BD-FACS 104 Mean CH # to 256 Relative CH#
Note on J50: Use this Table to convert BD-FACS Canto 105 Mean CH # to 384 Relative CH#